## Slide 1
EME Station Back End
Db6nt converter
DEM 144/28mhz transverter
Elecraft K3 reads 1296.000
P3 Panadapter
24” Monitor

## Slide 2: Microwave Power sensor Limitations
Doug Millar K6JEY
drzarkof56@yahoo.com

## Slide 3: Scope of Talk
RF power Sensor ratings
Test equipment used and useful
Clinic to measure DB attenuation on
  Power Meters
  Attenuators
  Spectrum Analyzers

## Slide 4: Background
Interest in metrology and the accuracy limits of RF power
Considerable experience with power meters
The arrival of a Boonton generator pushed me to explore the accuracy of sensors and their linearity.
I was only able to resolve four digits accurately with the calorimeter.
It turns out that is good enough.

## Slide 5: Definitions
Bolometer/thermistor- Resistance changes with exposure to RF
Thermocouple- Voltage is generated by exposure to RF in two dissimilar metals.
Thermopile- a series of dissimilar metals that generate DC voltage when exposed to DC or AC voltage.
Diode Sensor- generates a DC voltage through diode rectification.

## Slide 6: Traceability tools
At Boonton
NIST Micro Calorimeter for DC substitution
Wandel and Goltermann EPM 1 –
Accurate but little resolution
Trombone attenuator at 30MHz
Boonton 2520 .05db accuracy
At HP
0dbm thermistor reference sensor. (HP 478 h75)

## Slide 7: Micro calorimeter at NIST Calibration is done at 30-50Mhz

## Slide 8
Wandel and Goltermann EPM-1
1mw power meter

## Slide 9
Boonton 2520- 30MHz, -70 to +20 DBm.
Accuracy is .05db

## Slide 10
Lab standard at HP
Selected for low SWR. 10MHz to 500MHz
Used price $1,400 HP478-H75

## Slide 11: My Lab Traceability
DC source calibration- HP34401A DVM Calibrated to .0002%
DC Power source- Valhalla 2700b
Transfer standard- PRD 685 Calorimetric Thermopile (1mw= 100uv)
  Power meter - HP 8482B sensor on an HP E4418b. In Cal
  RF Power source- Boonton 2520 RF power calibrator. 30MHz +20to -70DBM
Accuracy .05DB or 1.2%

## Slide 12
PRD 670 Thermopile
.2236v (1mw) input =
.1160v output

## Slide 13: 2520 Meter Calibrator Output
Cal output is .7% to 1.2% accuracy
Based on Standards lab equipment at Boonton
  Thermopile DC substitution
  Calibrated trombone attenuator at 30MHz
  In Cal if .988mw to 1.012mw- 1.2%
Far exceeds the measurement accuracy.

## Slide 14: What are We Up against?
What is the list of uncertainties in measurement with a microwave power sensor?
What factors contribute the most?

## Slide 15: Sum of Uncertainties- Worst Case
The real problem -
--
3.8% mismatch?

## Slide 16
HP 432 and
8478
HP 8481A
HP 8481D
E4412A EPROM
Sensor
Power Range
Vs. Uncertainty

## Slide 17: Mismatch and Accuracy
The SWR of both the source and sensor have to be
Below 1.10:1 in order to add no more than .5% error to the reading. (20DB RL)
At 1.2 SWR the Inaccuracy is 1.5% (9DB RL)
Anything higher at either end due to any factor, and the inaccuracy skyrockets.
Sensor SWR and generator SWR are key to the accuracy problem.

## Slide 18
Gen HP 8481B 1.2db Anritsu 2442A- 1.02db
x
X= HP 436
And
HP8482A

## Slide 19: SensorSWR
Input coupling
Capacitor affects
Low freq. response

## Slide 20: Cal Factor/SWR vs Frequency for HP 18 GHz Sensor
SWR
Cal Factor
SWR
Cal Factor

## Slide 21: Accuracy Losses
0.5% Meter- fixed
4.5% Temperature- variable 1%?
1.8% Sensor Linearity- fixed
3.8% Mismatch loss- 1-2%
1.2% Calibrator calibration- .7% more likely
11.8% total uncertainty of reading worst case.
With some attention to detail, 6% is more like it.

## Slide 22: Perspective
Modern meters like the E4418b/E4412a
  Use averaging through DSP
  Have stored factors in the sensor
  Compensate for temperature
  Use computers to characterize both the meter and sensor.
Modern meters may constrain the errors to 5%.
Meters like the HP 436 may do a pretty good job.
Note- dropping the sensor or applying too much power, even for a second, can completely change the sensor linearity and accuracy

## Slide 23: Questions?
Thanks very much!

## Slide 24: Clinic
Use your score sheet to measure
  Power Sensors with your meter
  Spectrum Analyzer attenuator linearity
  Attenuator value at 30MHz

## Slide 25: Sensor Linearity Record .05db/1.2%

## Slide 26: Resources
HP Application Note on RF Power Measurement
64-1C
PRD Calorimeter Manuals 680 series.
Anritsu manual on 2400 power sensors
  https://www.atecorp.com/ATECorp/media/pdfs/data-sheets/Anritsu-MA2400-Series_Manual.pdf